## Tabelle1
Procedure for determining the specific activities of radionuclides in fish by gamma spectrometry
| G-g-SPEKT-FISCH-01 | [6 empty] | Version January 2016
Procedures manual for monitoring of radioactive substances in the environment and of external radiation (ISSN 1865-8725)
| Sample identification: |  |  |  | Variant 1: single-line radionuclide Cs-137
| #Number of parameters p: | 13 |  | Colors: | values from Vbasic
| k_alpha: quantile for (1-a) | 3 |  |  | Excel formulae (user)
| k_beta: quantile for (1-b) | 1.645 |  |  | Manual input of values
| gamma: probability for conf.limits | 0.05 |  |  | Def. Excel-Variable (user)
#Keywords
| Data input: |  | variable names: |  |  | Uncertainty budget:
| #Values of parameters p: | Unit: |  | Input values |  | partial | uncertainty | budget
[5 empty] | StdDev: | derivatives | budget: | in %
p1 | #Number of gross counts Ng: |  | Ng | 650.16 | 25.4982 | 0.00048129 | 0.012272 | 63.3142
p2 | trapezoidal BG count rate | 1/s | RT | 0.00271 | 0.0001403 | -34.653 | 0.004862 | 9.9377
p3 | b/2L |  | b2L | 0.523 | 0 | 0 | 0 | 0
p4 | detection efficiency |  | eps | 0.03056 | 0.00088624 | -7.1665 | 0.0063512 | 16.958
p5 | emission probability, 661,7 keV |  | p_gamm | 0.8499 | 0.001997 | -0.25769 | 0.0005146 | 0.11133
p6 | ash mass | kg | mA | 0.01317 | 5.268e-05 | -16.6292 | 0.00087603 | 0.32263
p7 | ratio fresh mass / ash mass |  | qF | 83.54 | 1.6708 | -0.0026216 | 0.0043801 | 8.0656
p8 | Time elapsed Sampling-->StartMeasum | s | tA | 19788000 |  | 1.6008e-10 | 0 | 0
p9 | half-live of Cs-137 | s | tCs137 | 948300000 |  | -3.3466e-12 | 0 | 0
p10 | self-attenuation correction factor |  | _f2 | 0.976 | 0.007808 | 0.22439 | 0.0017521 | 1.2905
p11 | counting duration | s | tm | 72000 | 0 | -4.346e-06 | 0 | 0
p12 | net count rate of the BG peak | 1/s | RnNE | 0 | 0 | -34.653 | 0 | 0
p13 | counting duration of the BG spectrum | s | tNE | 72000 | 0 | 0 | 0 | 0
(the list may be continued here)
| Model section |  | c = Factor * Rn
| auxiliary equations: |  |  | (formulae)
| #Gross counting rate Rg: | 1/s | Rg | 0.00903
h1 | decay correction measuring period |  | _f3 | 0.99997
h2 | decay correction Sampling-->StartMes |  | _f4 | 1.0146
(the list may be continued here)
(formulae)
| #Net counting rate Rn: | 1/s | Rn | 0.00632
| #Calibration factor, proc.dep.: | Bq*s/kg | Factor | 34.653
| #Value output quantity: | Bq/kg | Result | 0.21901 | 0.040542 | <-- output value modifiable by VBA
| #Combined standard uncertainty: | Bq/kg | uResult | 0.015423 |  |  |  |  | kombin. Stdmessunsicherheit
| #Decision threshold: | Bq/kg |  | 0.02489
| #Detection limit: | Bq/kg |  | 0.040542
further derived quantities (ISO 11929):
| auxiliary quantity omega |  | Omega | 1
| Best estimate | Bq/kg | BestWert | 0.21901
| Uncertainty of best estimate | Bq/kg |  | 0.015423
| lower confidence limit | Bq/kg |  | 0.18878 |  |  |  |  | u. Grenze d. Vertrauensbereichs
| upper confidence limit | Bq/kg |  | 0.24924 |  |  |  |  | o. Grenze d. Vertrauensbereichs
Calculate!